## CT Scanner
Vendor Name:
CT Scanner Model:
Mandatory Requirements * | Meet (insert YES or NO in cell) |  | Provide Technical Information
*Proposal will be disqualified if all Mandatory Requirements are not met. |  |  |  | YT ??'s
1 year manufacturer warranty, 5 year part & labor extended warranty, 6 year maintenance agreement that begins after completed installation included in purchase price (shown as a separate line item)
128 slice or higher or equivalent
3D analyzer
ECG Gating and motion correction
Calcium scoring package
Computer hardware to provide KCH with fastest image reconstruction available in seller's product line, including computer processor(s), disk drives, data acquisition, input/output units
Full DICOM 3.0 functionality
Full UPS, partial UPS is not desired.
Gating/LV with motion analysis
Most current in production, up-to-date model, first rate quality. All applicable options and software must be most current, up-to-date for the proposed model.
New, not refurbished or reconditioned
Optical disk drive or equivalent for archival storage-optional
Radiation reduction
Remote diagnotics
Slip ring gantry with detector assembly
Software for data acquisition, image reconstruction, image enhancement and image manipulation and image analysis
Software for image reconstruction in other planes (sagittal, coronal, parasagittal, paracoronal)
Software for scan image enlargement (prospective zoom)
Software for scan localization using reconstructed Anterior Projection (AP) and lateral scanned projection radiographs
Stent evaluation
X-ray generator
X-ray tube and housing assembly
Equipment/Computer hardware: | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Provide Technical Information | YT ??'s
- computer processor(s)
- data acquisition
- detectors |  |  |  | ??
- disk drives
- gantry, minimum 78cm - must be suitable for bariatric patients
- independent workstation, can be located up to 100m from the CT scanner
- input/output units
- operator's console, ergonomically designed, 2 monitors, truly multi-tasking with separate data basis for each
- patient couch, must adjust vertically electrically to accommodate physically challenged patients, minimum 500 lb capacity
- x-ray generator
- x-ray tube
CT phantom (complies with ACR accreditation standards)
Dual head contract injector, ceiling mount MEDRAD preferred, or pedistal style dual head injector
Computer software: | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Provide Technical Information | YT ??'s
- image reconstruction in other planes (sagittal, coronal, parasagittal, paracoronal)
2D/3D image acquisition
3D analizor
Ability to perform dynamic scanning for minimum 40 seconds.
Ability to see first slice before reconstruction
ECG Gating and motion correction
Automatic scan sequences including: the option to reconstruct & display images between scans, ability to easily make mid-study parameter changes and prospective gating to reduce raditiation exposure
Calcicum scoring package
Calcium score analysis
Calcium score reports
Cardiac reconstruction
CT Fluoroscopy, including ceiling mount, backlit 48" monitor |  |  |  | ?
Data acquisition in axial, dynamic and volumetric scanning modes
Gating/LV with motion analysis
High-contrast spatial resolution: 14.0 lp/cm or greater at MTT custoff for scan field of view for head scans, reconstructed and displayed over full matrix
Image analysis
Image display matric size: 1k H 1k or larger
Image enhancement
Image manipulation
Image reconstruction
Iterative Reconstruction, both 1st and 2nd generation
Low contrast spatial resolution: 3mm or smaller at 0.35% contrast at 4xGy or less for 10mm slice thickness
Maximum scan cycle time (scan, reconstruct, film, archive) of 10 seconds
Maximum scan field of view: 48 cm or greater |  |  |  | spelling?
Minimum rotational scan time: 0.6 sec (360E scan)
Radiation reduction
Scan localization using reconstructed AP and lateral scanned projection radiographs |  |  |  | spelling?
Stent evaluation
Cardiac Package: | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary) | YT ??'s
Able to perform ALL current cardiac studies (including scalcium scoring, Cardiac CTA, stent patency evaluation, LV volumes, etc.)
Cardiologist/Tech will be responsible for post processing
Neurological Package: | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary) | YT ??'s
Able to perform ALL current neurological studies (including perfusion studies, provide maps of MTT, CBF and CBV)
Neurologist/Tech will be responsible for post processing
Colonoscopy | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary) | YT ??'s
Include all software and hardware to perform a virtual colonoscopy
Radiologist/Tech will be responsible for post processing
CT Angiography, including run-offs | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary) | YT ??'s
Include all software and hardware to perform angiography
Must not require a test bolus for bolus tracking studies
Radiologist/Tech will be responsible for post processing
Orthopedic Imaging | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary) | YT ??'s
Include all software and hardware to perform orthopedic imaging
MAR
Delivery/Transporation Requirements | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Actual ship date from factory and container sail date must be communicated to KCH.
Equipment to be delivered on mutually acceptable date and time.
Equipment to remain in manufacturer's packaging from factory to KCH.
Freight charges to be included in pricing.
Rigging & coordination to be communicated with KCH.
Installation Requirements | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
All materials used will be new and of first-rate quality
All necessary trades will be provided and their costs included.
Equipment test protocol will be provided prior to installation
Work to be performed by bona-fide employees or authorized/certified distributor
Installation contractors will pull all applicable permits, including but not limited to: construction, electrical
Interface with IT, Imaging department, RIS/PACs (Fuji) and Radiologist connections at no additional cost
Must be fully insured to cover installation (at minimum, worker's comp, contractor's liability, property damage)
Provide all State and Federal regulatory documentation
Provide installation and pre-staging test results
Wall penetrations (if any) will be done with UL fire stopping materials
Within 30 days after award, provide drawing detailing all construction, electric and other work to be performed by vendor.
Work to be performed by bona-fide employees or authorized/certified distributor
Specify room requirements and site readiness to support the system. Drawings may be placed immediately behind this table.
Training Requirements - Two Phase | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Additional, optional training available if requested
American College of Cardiology (ACC) accredited
American Heart Associateion (AHA) accredited
ASRT accredited
Online training modules available to WHR techs throughout life of agreement at no additional charge.
On-site (KCH) training is required at no additional cost. Training will be for 3-4 techs. Due to staffing requirements, off-site training is not feasible for KCH.
YT ??'s
Phase one: | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary) | More detail?
Minimum 40 hours, on-site at KCH for both technologists and physicians
YT ??'s
Phase two: | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Minimum 72 hours, on-site at KCH for both technologists and physicians within one year of installation
Warranty Requirements | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Any part that is replaced during the warranty period will have its own full warranty that begins at its installation date
Guarantee availability of repair parts for minimum of 10 years after final equipment acceptance.
Min. response time of 1/2 hr via telephone, and onsite within 2 hrs if unit is down per telephone call. 4 hour on site response time if unit is still operational
Provide new replacement of any defective part that requires repair more than twice during the warranty period
Provide scheduled inspections and preventative maintenance |  |  |  | maximum or minimum?
Replace or repair any/all defective equipment and/or software at no cost to KCH (parts and labor), M-F, 6am - 10pm
System warranty will be extended for one month after each subsequent replacement
Replacement Tubes to be covered under warranty (parts and labor)
One (1) year minimum full manufacturer's warranty (parts & labor) included in purchase price.
Five (5) year extended warranty (parts & labor) included in purchase price - show as separate line item on proposal
Six (6) year full service maintenance agreement included in purchase price - show as a separate line item on proposal
Remote monitoring of system and tube guardian or equivalent
Support Requirements | Meet / Exceed (insert YES or NO in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Minimum response time of 1/2 hour via telephone, 2 hours on-site if system is down or four (4) hours on-site if system is still operational
Service technican to be located in the state of Hawaii throughout the life of the Service Agreement.
MONTHLY uptime quarantee of 98%. If uptime is not met, what compensation will be provided to KCH or KOH?
No separate licensing fee for use of software
Provide 1 set complete replacement parts list upon acceptance - hard copy + electronic
Provide 1 set of diagnostic software, manuals and phantoms upon acceptance - hard copy + electronic, if available
Provide 2 sets of Maintenance & Service Manuals upon acceptance - hard copies + electronic |  |  |  | separate out emergency repairs?
Supply 2 sets of Operator Manuals upon acceptance- hard copies + electronic
Provide updates to KCH when any applicable change to the manuals are made. Updates to be provided within 30 days of change. |  |  |  | what response time to you want?
Remote diagnostics at no charge
First and Second Support Years
Applications support
Documentation
In service training
Repair and replacement of defective equipment
Scheduled inspection/preventative maintenance
Software upgrades (see below details)
Hardware/software updating
All upgrades and any associated hardware to support them shall be brought to KCH's attention within 30 days of their release
Any upgrades or enhancements that add new capabilities to the system supplied will be available to HHSC for a minimum discount of 70% for 5 years.
No charge software upgrade, including parts, if enhancing existing capabilities
No charge software upgrade, not including parts, if adding new capabilities
No charge upgrades to make the CT comply with changing safety standards and regulations
General Questions | Expanded Answer
Describe your proposed system's ASIR features and capabilities, including dose reduction.
During what timeframe (Hawaii Standard Time) will software upgrades be conducted?
How does humity effect the system?
How will a large bore affect image resolution?
How will software upgrades be provided?
How will your proposed system address iterative image reconstruction?
Is there a separate licensing fee for use of your software? If yes, what is the fee and how often is it payable? Include in proposal as a separate line item.
List the system's scanning techniques available as standard options
State your proposed system's DICOM functionality and provide the DICOM conformance statement.
What discount will be given for any/all parts needed after warranty period expires?
What discount will upgrades and/or enhancements that add new capabilities or parameters to this equipment be extended and for how long?
What downtime is needed for scheduled preventative mantainance?
What is the expected life-span of the quoted equipment?
What is the lead time from receipt of p.o. of your quoted system?
What is the system's cooling system requirements (BTU/hr)?
What is the system's electrical requirements (V & kVA)?
What is the system's power consumption (kW)?
What other environmental requirements does the system have (humidity, etc.)?
What RH value does the room need?
What upgrades does your company anticipate coming to market in the next 3 - 5 years?
Cardiac Package:
What does your proposed package encompass?
Neurological Package:
What does your proposed package encompass?
Colonoscopy
What does your proposed package encompass?
CT Angiography, including run-offs
What does your proposed package encompass?
Orthopedic Imaging
Include all software and hardware to perform orthopedic imaging
Miscellaneous
For replacement parts that are not stored in Oahu, where are they stored and how long will they take to reach KCH? Is this guaranteed?
How many systems, same as you are proposing, have been installed in Hawaii for calendar year 2024 and 2025, up to the date of RFP release?
How many systems, same as you are proposing, have been installed in the U.S. for calendar years 2024 and 2025, up to the date of RFP release?
What holidays and other observances will be charged at overtime rate?
What is the anticipated downtime timeframe for installation? Is this guaranteed?
What is your definition of uptime and how do you calculate it?
What replacement parts are stored on Oahu?
What replacement parts are stored on the Big Island?
Where are the tubes stored?
What marketing tools are available from your company?
Final questions:
What is the closest location the proposed equipment is installed should we choose to perform a site-visit?
Why should you be awarded the contract for our new CT?
Any additional Information you would like us to know?

## Fluoroscopic - KCH Only
Vendor Name:
Fluoroscopic System Model Name/Number:
Mandatory Requirements * | Meet (insert YES or NO in cell) |  | Provide Technical Information
*Proposal will be disqualified if all Mandatory Requirements are not met. |  |  |  | YT ??'s
1 year manufacturer warranty, 5 year part & labor extended warranty, 6 year (warranty + ext warranty) maintenance agreement that begins after completed installation included in purchase price (shown as a separate line item)
Small footprint
System to include at a minimum: table, Bucky and grid system, digital detector, x-ray generator, x-ray tube and suspension system and collinator.
High patient weight limit (minimum 600lbs)
Bedside and remote operation
TIMS for Video swallow studies
Stirrups for Hysterosalpingograms
DSA capability
Stitching capability
MAR (Metal Artifact Reduction)
Table height can be adjusted to facilitate patient transfer from a mobile stretcher or wheelchair.
Most current in production, up-to-date model, first rate quality. All applicable options and software must be most current, up-to-date for the proposed model.
New, not refurbished or reconditioned
Optical disk drive or equivalent for archival storage-optional
Radiation reduction
Remote diagnotics
Software for data acquisition, image reconstruction, image enhancement and image manipulation and image analysis
Software for image reconstruction in other planes (sagittal, coronal, parasagittal, paracoronal)
Software for scan image enlargement (prospective zoom)
high patient weight limit (minimum 600lbs)
Stent evaluation
X-ray generator
X-ray tube and housing assembly
General: | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Provide Technical Information | YT ??'s
What is the system's boot up time from a cold boot?
Does the system have a sleep mode? If yes, what is the boot up time from the sleep mode?
What is the preview (exposure to view) time?
With what methods can the sytem be turned on?
What is the overall footprint of the system (length, width and height with mast/boom fully retracted)?
Where is the system manufactured (list components seperately if necessary)?
Does the system have any storage (gloves, markers, papers, etc.)?
Date of FDA approval?
Describe your marketing support for community outreach.
Where is your account rep located?
Where are your service techs located?
What exams can your systems NOT perform?
Provide 3-5 reasons your recommended system is stronger than your competitors.
Equipment/Computer hardware: | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Provide Technical Information | YT ??'s
- computer processor(s)
- data acquisition
- detectors |  |  |  | ??
- disk drives
- independent workstation, can be located up to 100m from the xray
- input/output units
- operator's console, ergonomically designed, 2 monitors, truly multi-tasking with separate data basis for each
Table: | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Provide Technical Information | YT ??'s
Electromagnetic locking system
Electro-mechanical float top that moves in 4 directions
Electro-mechanical float top that tilts and moves in minimum of 6 directions |  |  |  | ??
High-frequency generator (no high-voltage cables)
Pedestal floor mount
table top hand control
Table height can be adjusted to facilitate patient transfer from a mobile stretcher or wheelchair.
Delivery/Transporation Requirements | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Actual ship date from factory and container sail date must be communicated to KCH in a timely manner.
Equipment to be delivered on mutually acceptable date and time.
Equipment to remain in manufacturer's packaging from factory to KCH.
Freight charges to be included in pricing.
Rigging & coordination to be communicated with KCH.
Installation Requirements | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
All materials used will be new and of first-rate quality
All necessary trades will be provided and their costs included.
Equipment test protocol will be provided prior to installation
Installation company will pull all applicable permits
Interface with IT, Imaging department, RIS/PACs (Fuji) and Radiologist connections at no additional cost
Must be fully insured to cover installation (at minimum, worker's comp, contractor's liability, property damage)
Provide all State and Federal regulatory documentation
Provide installation and pre-staging test results
Within 30 days after award, provide drawing detailing all construction, electric and other work to be performed by vendor.
Wall penetrations (if any) will be done with UL fire stopping materials
Specify room requirements and site readiness to support the system. Drawings may be place immediately behind this table.
Work to be performed by bona-fide employees or authorized/certified distributor
Preventative Maintenance | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Work to be performed in off-peak hours (11pm - 5am HST)
Document detailing work performed to be provided to KCH within 14 days of service.
Training Requirements - Two Phase | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Additional, optional training available if requested
American College of Cardiology (ACC) accredited
American Heart Associateion (AHA) accredited
ASRT accredited
Online training modules available to WHR techs throughout life of agreement at no additional charge.
On-site (KCH) training is required at no additional cost. Training will be for 3-4 techs. Due to staffing requirements, off-site training is not feasible for KCH.
YT ??'s
Phase one: | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary) | More detail?
Minimum 40 hours, on-site at KCH for both technologists and physicians
YT ??'s
Phase two: | Meet / Exceed (insert word in cell) | Does Not / Can Not Meet | Comments (if necessary)
Minimum 72 hours, on-site at KCH for both technologists and physicians within one year of installation
Warranty Requirements | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Any part that is replaced during the warranty period will have its own full warranty that begins at equipment's acceptance date
Guarantee availability of repair parts for minimum of 10 years after final equipment acceptance.
Min. response time of 1/2 hr via telephone, and onsite within 2 hrs if unit is down per telephone call. 4 hour on site response time if unit is still operational
Provide new replacement of any defective part that requires repair more than twice during the warranty period
Provide scheduled inspections and preventative maintenance |  |  |  | maximum or minimum?
Replace or repair any/all defective equipment and/or software at no cost to KCH (parts and labor), M-F, 8am - 10pm HST
System warranty will be extended for one month after each subsequent replacement
Tubes to be covered under warranty (parts and labor)
One (1) year minimum full manufacturer's warranty (parts & labor) included in purchase price.
Five (5) year extended warranty (parts & labor) included in purchase price - show as separate line item on proposal
Six (6) year full service maintenance agreement included in purchase price - show as a separate line item on proposal. If preventative maintenance is included in the first year, state as such.
Support Requirements | Meet / Exceed (insert word in cell) | Does not Meet (insert word in cell) | Comments (if necessary)
MONTHLY uptime quarantee of 98%. If uptime is not met, what compensation will be provided to KCH or KOH?
Service technican to be located in the state of Hawaii throughout the life of the Agreement.
No separate licensing fee for use of software
Provide 1 set complete replacement parts list upon acceptance - hard copy + electronic
Provide 1 set of diagnostic software, manuals and phantoms upon acceptance - hard copy + electronic |  |  |  | separate out emergency repairs?
Provide 2 sets of Maintenance & Service Manuals upon acceptance - hard copies + electronic
Supply 2 sets of Operator Manuals upon acceptance- hard copies + electronic |  |  |  | what response time to you want?
Provide updates to KCH when any applicable change to the manuals are made. Updates to be provided within 30 days of change.
Remote diagnostics at no charge
First and Second Support Years
Applications support
Documentation
In service training
Repair and replacement of defective equipment
Scheduled inspection/preventative maintenance
Software upgrades (see below details)
Hardware/software updating
All upgrades and any associated hardware to support them shall be brought to KCH's attention within 30 days of their release
Any upgrades or enhancements that add new capabilities to the system supplied will be available to HHSC for a minimum discount of 70% for 5 years.
No charge software upgrade, including parts, if enhancing existing capabilities
No charge software upgrade, not including parts, if adding new capabilities
No charge upgrades to make the digital x-ray unit comply with changing safety standards and regulations
General Questions | Answers
Describe your proposed system's features and capabilities, including dose reduction.
During what timeframe (Hawaii Standard Time) will software upgrades be conducted?
How does humity effect the system?
What RH value does the room need?
How will software upgrades be provided?
Is there a separate licensing fee for use of your software? If yes, what is the fee and how often is it payable? Include in proposal as a separate line item.
List the system's scanning techniques available as standard options
State your proposed system's DICOM functionality and provide the DICOM conformance statement.
What discount will be given for any/all parts needed after warranty period expires?
What discount will upgrades and/or enhancements that add new capabilities or parameters to this equipment be extended and for how long?
What downtime is needed for scheduled preventative mantainance?
What is the expected life-span of the quoted equipment?
What is the lead time from receipt of p.o. of your quoted system?
What is the system's cooling system requirements (BTU/hr)?
What is the system's electrical requirements (V & kVA)?
What is the system's power consumption (kW)?
What other environmental requirements does the system have (humidity, etc.)?
What upgrades does your company anticipate coming to market in the next 3 - 5 years?
Stitching
What does your proposed package encompass?
Orthopedic Imaging
Include all software and hardware to perform orthopedic imaging
Miscellaneous
For replacement parts that are not stored in Oahu, where are they stored and how long will they take to reach KCH? Is this guaranteed?
How many systems, same as you are proposing, have been installed in Hawaii for calendar year 2024 and 2025, up to the date of RFP release?
How many systems, same as you are proposing, have been installed in the U.S. for calendar years 2024 and 2025, up to the date of RFP release?
What holidays and other observances will be charged at overtime rate?
What is the anticipated downtime timeframe for installation? Is this guaranteed?
What is your definition of uptime and how do you calculate it?
If agreed upon uptime is not met, what compensation will you provide?
What replacement parts are stored on Oahu?
What replacement parts are stored on the Big Island?
What do you plan to do with our existing equipment?
Final questions:
What is the closest location the proposed equipment is installed should we choose to perform a site-visit?
Why should you be awarded the contract for our new fluoroscopic system?
Any additional Information you would like us to know?

## XRay
Vendor Name:
Diagnostic X-ray Model Name/Number:
Mandatory Requirements * | Meet (insert YES or NO in cell) |  | Provide Technical Information
*Proposal will be disqualified if all Mandatory Requirements are not met. |  |  |  | YT ??'s
1 year manufacturer warranty, 5 year part & labor extended warranty, 6 year (warranty + ext warranty) maintenance agreement that begins after completed installation included in purchase price (shown as a separate line item)
Small footprint
System to include at a minimum: table, Bucky and grid system, digital detector, x-ray generator, x-ray tube and suspension system and collinator.
High patient weight limit (minimum 600lbs)
Stitching capability
MAR (Metal Artifact Reduction)
Table height can be adjusted to facilitate patient transfer from a mobile stretcher or wheelchair.
Most current in production, up-to-date model, first rate quality. All applicable options and software must be most current, up-to-date for the proposed model.
New, not refurbished or reconditioned
Optical disk drive or equivalent for archival storage-optional
Radiation reduction
Remote diagnotics
Software for data acquisition, image reconstruction, image enhancement and image manipulation and image analysis
Software for image reconstruction in other planes (sagittal, coronal, parasagittal, paracoronal)
Software for scan image enlargement (prospective zoom)
Software for scan localization using reconstructed Anterior Projection (AP) and lateral scanned projection radiographs
Stent evaluation
X-ray generator
X-ray tube and housing assembly
General: | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Provide Technical Information | YT ??'s
What is the system's boot up time from a cold boot?
Does the system have a sleep mode? If yes, what is the boot up time from the sleep mode?
What is the preview (exposure to view) time?
With what methods can the sytem be turned on?
What is the overall footprint of the system (length, width and height with mast/boom fully retracted)?
Where is the system manufactured (list components seperately if necessary)?
Does the system have any storage (gloves, markers, papers, etc.)?
Date of FDA approval?
Describe your marketing support for community outreach.
Where is your account rep located?
Where are your service techs located?
What exams can your system NOT perform?
Provide 3-5 reasons your recommended system is stronger than your competitors.
Equipment/Computer hardware: | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Provide Technical Information | YT ??'s
- computer processor(s)
- data acquisition
- detectors |  |  |  | ??
- disk drives
- independent workstation, can be located up to 100m from the xray
- input/output units
- operator's console, ergonomically designed, 2 monitors, truly multi-tasking with separate data basis for each
Table: | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Provide Technical Information | YT ??'s
Electromagnetic locking system
Electro-mechanical float top that moves in 4 directions
Electro-mechanical float top that tilts and moves in minimum of 6 directions |  |  |  | ??
High-frequency generator (no high-voltage cables)
Pedestal floor mount
table top hand control
Delivery/Transporation Requirements | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Actual ship date from factory and container sail date must be communicated to KCH in a timely manner.
Equipment to be delivered on mutually acceptable date and time.
Equipment to remain in manufacturer's packaging from factory to KCH.
Freight charges to be included in pricing.
Rigging & coordination to be communicated with KCH.
Installation Requirements | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
All materials used will be new and of first-rate quality
All necessary trades will be provided and their costs included.
Equipment test protocol will be provided prior to installation
Installation company will pull all applicable permits
Interface with IT, Imaging department, RIS/PACs (Fuji) and Radiologist connections at no additional cost
Must be fully insured to cover installation (at minimum, worker's comp, contractor's liability, property damage)
Provide all State and Federal regulatory documentation
Provide installation and pre-staging test results
Within 30 days after award, provide drawing detailing all construction, electric and other work to be performed by vendor.
Wall penetrations (if any) will be done with UL fire stopping materials
Work to be performed by bona-fide employees or authorized/certified distributor
Specify room requirements and site readiness to support the system. Drawings may be placed immediately behind this table.
Preventative Maintenance | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Work to be performed in off-peak hours (11pm - 5am)
Document detailing work performed to be provided to KCH within 14 days of service.
Training Requirements - Two Phase | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Additional, optional training available if requested
American College of Cardiology (ACC) accredited
American Heart Associateion (AHA) accredited
ASRT accredited
Online training modules available to WHR techs throughout life of agreement at no additional charge.
On-site (KCH) training is required at no additional cost. Training will be for 3-4 techs. Due to staffing requirements, off-site training is not feasible for KCH.
YT ??'s
Phase one: | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary) | More detail?
Minimum 40 hours, on-site at KCH for both technologists and physicians
YT ??'s
Phase two: | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Minimum 72 hours, on-site at KCH for both technologists and physicians within one year of installation
Warranty Requirements | Meet / Exceed (insert word in cell) | Does Not Meet (insert words in cell) | Comments (if necessary)
Any part that is replaced during the warranty period will have its own full warranty that begins at equipment's acceptance date
Guarantee availability of repair parts for minimum of 10 years after final equipment acceptance.
Min. response time of 1/2 hr via telephone, and onsite within 2 hrs if unit is down per telephone call. 4 hour on site response time if unit is still operational
Provide new replacement of any defective part that requires repair more than twice during the warranty period
Provide scheduled inspections and preventative maintenance |  |  |  | maximum or minimum?
Replace or repair any/all defective equipment and/or software at no cost to KCH (parts and labor), M-F, 8am - 10pm
System warranty will be extended for one month after each subsequent replacement
Tubes to be covered under warranty (parts and labor)
One (1) year minimum full manufacturer's warranty (parts & labor) included in purchase price.
Five (5) year extended warranty (parts & labor) included in purchase price - show as separate line item on proposal
Six (6) year full service maintenance agreement included in purchase price - show as a separate line item on proposal
Support Requirements | Meet / Exceed (insert words in cell) | Does not Meet (insert words in cell) | Comments (if necessary)
MONTHLY uptime quarantee of 98%. If uptime is not met, what compensation will be provided to KCH or KOH?
Service technican to be located in the state of Hawaii throughout the life of the Agreement.
No separate licensing fee for use of software
Provide 1 set complete replacement parts list upon acceptance - hard copy + electronic
Provide 1 set of diagnostic software, manuals and phantoms upon acceptance - hard copy + electronic |  |  |  | separate out emergency repairs?
Provide 2 sets of Maintenance & Service Manuals upon acceptance - hard copies + electronic
Supply 2 sets of Operator Manuals upon acceptance- hard copies + electronic |  |  |  | what response time to you want?
Provide updates to KCH when any applicable change to the manuals are made. Updates to be provided within 30 days of change.
Remote diagnostics at no charge
First and Second Support Years
Applications support
Documentation
In service training
Repair and replacement of defective equipment
Scheduled inspection/preventative maintenance
Software upgrades (see below details)
Hardware/software updating
All upgrades and any associated hardware to support them shall be brought to KCH's attention within 30 days of their release
Any upgrades or enhancements that add new capabilities to the system supplied will be available to HHSC for a minimum discount of 70% for 5 years.
No charge software upgrade, including parts, if enhancing existing capabilities
No charge software upgrade, not including parts, if adding new capabilities
No charge upgrades to make the digital x-ray unit comply with changing safety standards and regulations
General Questions | Expanded Answer
Describe your proposed system's features and capabilities, including dose reduction.
During what timeframe (Hawaii Standard Time) will software upgrades be conducted?
How does humity effect the system?
How will software upgrades be provided?
Is there a separate licensing fee for use of your software? If yes, what is the fee and how often is it payable? Include in proposal as a separate line item.
List the system's scanning techniques available as standard options
State your proposed system's DICOM functionality and provide the DICOM conformance statement.
What discount will be given for any/all parts needed after warranty period expires?
What discount will upgrades and/or enhancements that add new capabilities or parameters to this equipment be extended and for how long?
What downtime is needed for scheduled preventative mantainance?
What is the expected life-span of the quoted equipment?
What is the lead time from receipt of p.o. of your quoted system?
What is the system's cooling system requirements (BTU/hr)?
What is the system's electrical requirements (V & kVA)?
What is the system's power consumption (kW)?
What other environmental requirements does the system have (humidity, etc.)?
What RH value does the room need?
What upgrades does your company anticipate coming to market in the next 3 - 5 years?
Stitching
What does your proposed package encompass?
Orthopedic Imaging
Include all software and hardware to perform orthopedic imaging
Miscellaneous
For replacement parts that are not stored in Oahu, where are they stored and how long will they take to reach KCH? Is this guaranteed?
How many systems, same as you are proposing, have been installed in Hawaii for calendar year 2024 and 2025, up to the date of RFP release?
How many systems, same as you are proposing, have been installed in the U.S. for calendar years 2024 and 2025, up to the date of RFP release?
What holidays and other observances will be charged at overtime rate?
What is the anticipated downtime timeframe for installation? Is this guaranteed?
What is your definition of uptime and how do you calculate it?
What replacement parts are stored on Oahu?
What replacement parts are stored on the Big Island?
Where are the tubes stored?
What do you plan to do with our existing equipment that we trade in?
Final questions:
What is the closest location the proposed equipment is installed should we choose to perform a site-visit?
Why should you be awarded the contract for our new digital x-ray unit?
Any additional Information you would like us to know?